(FACS Calibur, Cell Quest program)

Adjustment of samples:

GFP:

- Transfer each 30μl(appropriate doses)/1ml of cell suspension into 1.5ml tubes (GFP and hCD4).

- Re-suspended cells was centrifuged at 4000rpm for 1min and aspirate the supernatant.

Cells were re-suspended with 1ml of PBS or FACS buffer containing 1μl of PI(Stock:2mg/ml→2μg/ml(1000-fold dilution)) for FACS analysis.

- Filtrate to a FACS tube with membrane.

- Vortex it before use.

- It can be stored at 4℃.

hCD4:

- 30μl(appropriate doses) of re-suspended cells was centrifuged at 4000rpm for 1min and aspirate the supernatant.

- Add 3μl of PE-hCD4 and the cells were incubated on ice for 20min in the dark.

- 1ml of PBS was added into the mixture and centrifuge at 4000rpm for 1min.

- Supernatant was discarded.

- Cells were re-suspended with 1ml of PBS or FACS buffer containing 1μl of PI(1:1000dilution) for FACS analysis.

- Filtrate to a FACS tube.

- Fluorescence analysis by a FACS calibur.

Prepare for control GFP(－)PI(－) and GFP(＋)PI(－), too.

Fluorescence wavelength:

FL1 530－FITC, GFP

FL2 585－PE, Cy3

FL3 650－PI

FL4－APC (only in the case of checking four color checkbox)

Password: BDIS

Open [Cell Quest].

STATAS: Stanby

Command

＋B [Connect to Cytometer]: connect the caliber with the computer

＋1 [Detecters/Amps]

＋2 [Thereshold]

＋3 [Compensation]

＋4 [Status]

[Detecters/Amps]

This window is used to adjust the Voltage and Amplifiers Gain of the detectors so that the signals appear appropriately on the data plots.The light signals are generated by particles passing through the laser beam in the flow cytometer. These light signals are converted to electronic signals (voltages), and then assigned a channel number on a data plot. By adjusting the detectors and amplifyiers, you control where these signals appear on the dot plot.

Forward Scatter (FSC/P1):

This parameter is detected using a photdiode and is a measure of cell size. FSC can adjusted using the Voltage and Amp Gain. Logarithmic increments of E01, E02, and E03 increase the signal of the small events; E-1 reduces the signal of large events. For most applications the voltage is set to E-1 or E01.

Side Scatter (SSC/P2):

This parameter is detected using a photomultiplier tube (PMT) and is a measure of cell complexity. This signal is amplified by applying a voltage to the PMT. As the voltage is increased, the detector sensitivity increases.

FL1(P3), FL2(P4), and FL3(P5):

These parameters are also detected by PMTs and are a measure of fluorescent intensity. These parameters are amplified by applying a voltage to the PMTs.

Mode:

Log amplifyication is ofter used to analyze samples with a large dynamic range of fluorescence signals. For most applications, such as phenolyping the fluorescence parameters are to log. Linear amplification is usually used for light scatter parameters. It is also used for DNA cell cycle analysis.

[Threshold]:

When running immunity, the threshold is set on FSC to eliminate events such as debris. The default setting is 52.

[Compensation]

This allows you to adjust for spectral overlap when samples are stained detector used for another fluorochrome. For example, FITC appears primarily in the FL detector, but some of the fluorescence overlaps into the FL2 detector. PE appears primarily in the FL2 detector, but some of the fluorescence overlaps into FL1 and FL3 detectors. Compensation allows you to eliminate the spectral overlap electronically. Be sure all are set to 0.0 to start.

From the [Acquire] drop down menu select.

[Parameter Description]

In this window you will need to change the Folder, File ,Sample ID, and Parameter Settings.

Change Folder.:year, month then name your folder. Be sure to select your folder.

Change File.Select Change and rename the Custm Pefix.Remember, if you need to reacquire data that has already been saved, you will need to change the gile count to overwrite the old data. Select OK.

Name the parameters which correspond to different channels (refer to Dectors/Amps).

[Counters]

Expand with counter window by selecting the upper right box and change the Rate to Accumulate. This will allow you to see what is actually being collected.

[Acquisision and Storage]

In this window you can change the number of events to be counted and the gate. 10,000 to 20,000 cells is standard depending on the experiment.

Vortex the sample.

Press [Low] and [Run].

| Mode | Rate |
| --- | --- |
| LO | 12μl/min |
| MED | 35μl/min |
| HI | 60μl/min |

If you take a rest, press [STANDBY].

[Acquisition control] →check [Set up] in order not to save the data into hard disk→[Acquire]

Create Plots:

Use [Toolbar] on left to draw plots.

Click [Dotplot]

[Edit] →[Duplicate]

Dot plot

Contour plot 3DHistogramDensity plot

If you want to use the “FL4-Hight”, tick “Four Color” checkbox.

Pop-up menu listing the parameters. The cursor changes to the pop-up cursor

whenever it is in the x- or y-axis area.

First plot should always be FSC vs SSC

While your control is running, adjust the FSC amp gain and voltage (E00 or E-1) and SSC amp gain so that the population of interest is clearly visible towards the center of FSC vs SSC dot plot. Make sure FSC and SSC are in the linear (lin) mode. Notice how the dots move to the right of the display. You have amplified your signal tenfold. The light signals from the cells can be multiplied by thesettings below.

• E00–multiplies the signal by 100 or 1

• E01–multiplies the signal by 101 or 10

• E02–multiplies the signal by 102 or 100

• E03–multiplies the signal by 103 or 1000

• E-1–multiplies the signal by 10–1 or 0.1

E01, E02, and E03 are useful for increasing the signal of small events.

E-1 is useful for reducing the signal of large events.

Make sure you return the settings to E00 before you proceed.

[Detecter Amps]

SSC TOO LOW

FSC TOO LOW

Change from “Lin”to “Log” on Mode of P3, P4 and P5.

○Control GFP(－)PI(－)：setting of gate1

You may draw a gate around the population of interest to remove unwanted debris and aggregates. Do so by selecting a circle or polygon from the Tool Palette.

Region tool

①Rectangular

②Polygonal

③Histograms

④Elliptical

Select the high-density region.

[Gate list]→Gate:”G1=R1”

Ex. multicolor gating

Select Quadrant Marker tool  in the tool palette.

Drag the handle of the markers so the cross-hairs are at 101 and 101 on the x- and y-axes, respectively.

If necessary, adjust the FL1 and FL2 voltages in the Detector/Amps window to place the negative population in the lower-left quadrant.

Adjustment of voltage

Dotplot of FL1-H FL2-H

Dotplot of FL1-H FL3-H

Make the dotplots placed within the less than 10^1.

○GFP(－)PI(＋)：Removal of positive region, setting of gate 2

Dotplot of counts/FL3-H

Dotplot window→Gate: ”G1=R1”

Select Histogram-Marker tool . from the tool palette.

Select the negative region.

[Gates]→[Gate list]→Gate”G2=R1＊R2”

○GFP(＋)PI(－): Adjustment of Compensation

Pur on your first compensateon control (Ex. GFP). You may notice that the GFP signal fluoresces mostly in the FL1 channel, but bleeds over into the FL2 channel. This needs to be compensated. Goto the compensateon window that you opened previously and adjust the equation F2-  % FL1. View FL1 vs FL2 dot plot. As you increase the S% compensation watch as the GFP population is stracted from the FL2 channel and Placed solely in the single positive quadrant. Be careful not to over compensate. The median of the GFP population should equal that of the unstained population. Once the compensateon is somplete do not change the volatags for subsequent samples. Doing so will unbalance the compensations and the processs will have to be repeated.

Dotplot of FL1-H FL2-H

Dotplot window→Gate” G2=R1＊R2”

Properly compensated

Undercompensated

Overcompensated

[Gates]→[Gate list]→Gate”G3=R1＊R2＊R3”

Measure:

[Acquisition & storage]→[Collection criteria]

Acquisition will stop when “10000” of “G3= R1＊R2＊R3” events are counted.

Set the save site.

Once you are sure that your settings are correct you may start acquision. Uncheck [Set up] so as to save the data into hard disk. Click [Acquire].

When the computer beeps and the counters stop, the data file is sutomatically saved to your folder. Check to be sure each file number(.001 at end of file name) corresponds to your sample #. To save the data file before it has reached the total event count (if event rate is low and will not reach total) click [Pause] then [Save] in the [Acquisition Control].

[Stats]→[Quadrant Stats]

Save the settings:

Next time, when you want to refer the saved instrument settings, Select [Open] and navigate to your file then select SET and DONE. Tip: Your old data files (FCS files) contain the instrument setting used to generate them. You may choose an FCS file that has similar settings for the experiment that you are running.

Print:

Command＋A and Command ＋P

Command＋{ }→To select the sample No.

Save your analysis experiment document.

[Edit]→[Save As]

Next time, you can open the document as a previous settings of analysis experiment.

Troubleshooting Guide:

If you open up CellQuest, and cannot “Connect to the Cytometer”:

• Make sure the cytometer is turned on

• Quit CellQuest and restart the computer

If you do not see any data:

• Be sure voltages are correct

• Put on another sample

• Put on beads or something we know fluoresces (chicken red blood cells)

If data looks funny:

• Be sure compensation was set correctly (refer to compensation guide)

• Be sure you are adjusting correct channel (FL1 or FL2)

• Check gates – were they applied to plots? Click on plot and under plots dropdown click format dot plot

• Put on another sample and see how it looks (maybe all cells are dead)

If sample is acquiring very slowly:

• Be sure to vortex

• Check for clumps (If you can see a clump, then it needs to be filtered – we have 50 um filters)

• Change speed to HI